## OPEN (2) (hidden)
FLASH DR | [5 empty] | Conventional DR
Scattering foil: | 0 | um | (3 us pulse) |  |  | Scattering foil: | 500 | um | (0.5 us pulse)
FWHM | 3 | cm |  |  |  | FWHM | 17.3 | cm
dmax | 1 | cm |  |  |  | dmax | 0.8 | cm
d90 | 2.07 | cm |  |  |  | d90 | 2.14 | cm
d80 | 2.54 | cm |  |  |  | d80 | 2.5 | cm
d50 | 3.56 | cm |  |  |  | d50 | 3.18 | cm
Rp | 5.03 | cm |  |  |  | Rp | 4.14 | cm
Max Sample size | 1 | cm | < 8 % variation across field |  |  | Max Sample size | 6.07 | cm | < 9 % variation across field
| 1.8 | cm | < 23% variation across field |  |  |  | 10.4 | cm | < 22% variation across field
| 2.4 | cm | < 36% variation across field |  |  |  | 13.9 | cm | < 36% variation across field
pulse width | Dose / pulse (Gy) | Uncertainty (Gy) | Average DR (Gy/s) | integrated charge / pulse (x10-7 C) | Uncertainty (x10-7 C) | pulse width | Dose / pulse (Gy) | Uncertainty (Gy) | Average DR (Gy/s) | integrated charge / pulse (C) (x10-9) | Uncertainty (x10-9 C)
0.5 | 26.0657 | 1.1819 | 7819.7045 | 1.0098 | 0.0047183 | 0.5 | 0.050808 | 5.6222e-06 | 0.50808 | 9.2911 | 0.362
1 | 50.797 | 4.7129 | 15239.1035 | 2.0998 | 0.0073664
2 | 99.4025 | 4.1267 | 29820.7458 | 4.2263 | 0.0099058
3 | 148.1676 | 3.193 | 44450.2938 | 6.335 | 0.01116
4 | 183.5816 | 16.2832 | 55074.4723 | 8.546 | 0.011181
5 | 226.7893 | 8.5212 | 68036.776 | 10.7 | 0.0091493
6 | 257.4462 | 11.8697 | 77233.8652 | 12.5067 | 0.0050248
Scattering foil: | 100 | um | (0.5 us pulse)
FWHM | 7.6 | cm
dmax | 1.2 | cm
d90 | 2.88 | cm
d80 | 3.27 | cm
d50 | 4.04 | cm
Rp | 5.12 | cm
Max Sample size | 2.7 | cm | < 6 % variation across field
| 4.6 | cm | < 23% variation across field
| 6.1 | cm | < 36% variation across field
pulse width | Dose / pulse (Gy) | Uncertainty (Gy) | Average DR (Gy/s) | integrated charge / pulse (x10-7 C) | Uncertainty (x10-7 C)
0.5 | 3.0795 | 0.25071 | 923.8611 | 1.0098 | 0.0047183
1 | 6.1591 | 0.50142 | 1847.7222 | 2.0998 | 0.0073664
2 | 12.3181 | 1.0028 | 3695.4444 | 4.2263 | 0.0099058
3 | 18.4772 | 1.5043 | 5543.1667 | 6.335 | 0.01116
4 | 24.6363 | 2.0057 | 7390.8889 | 8.546 | 0.011181
5 | 30.7954 | 2.5071 | 9238.6111 | 10.7 | 0.0091493
6 | 36.9544 | 3.0085 | 11086.3333 | 12.5067 | 0.0050248
Scattering foil: | 200 | um | (0.5 us pulse)
FWHM | 11.2 | cm
dmax | 1.6 | cm
d90 | 2.76 | cm
d80 | 3.11 | cm
d50 | 3.81 | cm
Rp | 4.79 | cm
Max Sample size | 3.9 | cm | < 8 % variation across field
| 6.7 | cm | < 23% variation across field
| 8.9 | cm | < 36% variation across field
pulse width | Dose / pulse (Gy) | Uncertainty (Gy) | Average DR (Gy/s) | integrated charge / pulse (x10-7 C) | Uncertainty (x10-7 C)
0.5 | 1.5222 | 0.15841 | 456.6611 | 1.0098 | 0.0047183
1 | 3.0444 | 0.31682 | 913.3223 | 2.0998 | 0.0073664
2 | 6.0888 | 0.63365 | 1826.6445 | 4.2263 | 0.0099058
3 | 9.1332 | 0.95047 | 2739.9668 | 6.335 | 0.01116
4 | 12.1776 | 1.2673 | 3653.289 | 8.546 | 0.011181
5 | 15.222 | 1.5841 | 4566.6113 | 10.7 | 0.0091493
6 | 18.2664 | 1.9009 | 5479.9335 | 12.5067 | 0.0050248
Scattering foil: | 500 | um | (0.5 us pulse)
FWHM | 18 | cm
dmax | 1.2 | cm
d90 | 2.29 | cm
d80 | 2.67 | cm
d50 | 3.36 | cm
Rp | 4.35 | cm
Max Sample size | 6.3 | cm |  | < 8 % variation across field
| 10.8 | cm |  | < 22% variation across field
| 14.4 | cm |  | < 35% variation across field
pulse width | Dose / pulse (Gy) | Uncertainty (Gy) | Average DR (Gy/s) | integrated charge / pulse (x10-7 C) | Uncertainty (x10-7 C)
0.5 | 0.64719 | 0.051421 | 194.1569 | 1.0098 | 0.0047183
1 | 1.2944 | 0.10284 | 388.3139 | 2.0998 | 0.0073664
2 | 2.5888 | 0.20568 | 776.6278 | 4.2263 | 0.0099058
3 | 3.8831 | 0.30853 | 1164.9417 | 6.335 | 0.01116
4 | 5.1775 | 0.41137 | 1553.2556 | 8.546 | 0.011181
5 | 6.4719 | 0.51421 | 1941.5695 | 10.7 | 0.0091493
6 | 7.7663 | 0.61705 | 2329.8834 | 12.5067 | 0.0050248

## Open
FLASH DR | [5 empty] | Conventional DR
Scattering foil: | 0 | um | (3 us pulse) |  |  | Scattering foil: | 500 | um | (0.5 us pulse)
FWHM | 3 | cm |  |  |  | FWHM | 17.3 | cm
dmax | 1 | cm |  |  |  | dmax | 0.8 | cm
d90 | 2.07 | cm |  |  |  | d90 | 2.14 | cm
d80 | 2.54 | cm |  |  |  | d80 | 2.5 | cm
d50 | 3.56 | cm |  |  |  | d50 | 3.18 | cm
Rp | 5.03 | cm |  |  |  | Rp | 4.14 | cm
Max Sample size | 1 | cm | < 8 % variation across field |  |  | Max Sample size | 6.07 | cm | < 9 % variation across field
| 1.8 | cm | < 23% variation across field |  |  |  | 10.4 | cm | < 22% variation across field
| 2.4 | cm | < 36% variation across field |  |  |  | 13.9 | cm | < 36% variation across field
pulse width | Dose / pulse (Gy) | Uncertainty (%) | Average DR (Gy/s) | Integrated Charge / pulse (C) | Uncertainty (%) | pulse width | Dose / pulse (Gy) | Uncertainty (Gy) | Average DR (Gy/s) | Integrated Charge / pulse (C) | Uncertainty (%)
0.5 | 24.3225 | 2.635% | 7296.75 | 1.0043e-07 | 0.7697% | 0.5 | 0.050808 | 5.6222e-06 | 0.50808 | 9.3034e-09 | 3.822%
1 | 43.885 | 2.623% | 13165.5 | 2.068e-07 | 0.2163%
2 | 93.42 | 3.61% | 28026 | 4.1826e-07 | 0.09148%
3 | 139.46 | 1.05% | 41838 | 6.3128e-07 | 0.1497%
4 | 172.33 | 3.89% | 51699 | 8.4536e-07 | 0.2066%
5 | 226.69 | 4.23% | 68007 | 1.0584e-06 | 0.1134%
6 | 259.11 | 2.25% | 77733 | 1.2442e-06 | 0.06015%
Scattering foil: | 100 | um | (0.5 us pulse)
FWHM | 7.6 | cm
dmax | 1.2 | cm
d90 | 2.88 | cm
d80 | 3.27 | cm
d50 | 4.04 | cm
Rp | 5.12 | cm
Max Sample size | 2.7 | cm | < 6 % variation across field
| 4.6 | cm | < 23% variation across field
| 6.1 | cm | < 36% variation across field
pulse width | Dose / pulse (Gy) | Uncertainty (Gy) | Average DR (Gy/s) | Integrated Charge / pulse (C) | Uncertainty (x10-7 C)
0.5 | 3.0534 | 4.21% | 916.0242 | 1.0043e-07 | 0.7697%
1 | 5.5468 | 4.05% | 1664.0524 | 2.068e-07 | 0.2163%
2 | 11.9112 | 4.501% | 3573.3743 | 4.1826e-07 | 0.09148%
3 | 18.073 | 2.522% | 5421.8953 | 6.3128e-07 | 0.1497%
4 | 22.7185 | 4.328% | 6815.5543 | 8.4536e-07 | 0.2066%
5 | 30.2908 | 4.489% | 9087.2508 | 1.0584e-06 | 0.1134%
6 | 35.1055 | 2.507% | 10531.6609 | 1.2442e-06 | 0.06015%
Scattering foil: | 200 | um | (0.5 us pulse)
FWHM | 11.2 | cm
dmax | 1.6 | cm
d90 | 2.76 | cm
d80 | 3.11 | cm
d50 | 3.81 | cm
Rp | 4.79 | cm
Max Sample size | 3.9 | cm | < 8 % variation across field
| 6.7 | cm | < 23% variation across field
| 8.9 | cm | < 36% variation across field
pulse width | Dose / pulse (Gy) | Uncertainty (Gy) | Average DR (Gy/s) | Integrated Charge / pulse (C) | Uncertainty (x10-7 C)
0.5 | 1.5666 | 4.21% | 469.9823 | 1.0043e-07 | 0.7697%
1 | 2.8636 | 4.05% | 859.0904 | 2.068e-07 | 0.2163%
2 | 6.1773 | 4.501% | 1853.203 | 4.1826e-07 | 0.09148%
3 | 9.3647 | 2.522% | 2809.4094 | 6.3128e-07 | 0.1497%
4 | 11.7838 | 4.328% | 3535.1453 | 8.4536e-07 | 0.2066%
5 | 15.788 | 4.489% | 4736.3898 | 1.0584e-06 | 0.1134%
6 | 18.3187 | 2.507% | 5495.5967 | 1.2442e-06 | 0.06015%
Scattering foil: | 500 | um | (0.5 us pulse)
FWHM | 18 | cm
dmax | 1.2 | cm
d90 | 2.29 | cm
d80 | 2.67 | cm
d50 | 3.36 | cm
Rp | 4.35 | cm
Max Sample size | 6.3 | cm |  | < 8 % variation across field
| 10.8 | cm |  | < 22% variation across field
| 14.4 | cm |  | < 35% variation across field
pulse width | Dose / pulse (Gy) | Uncertainty (Gy) | Average DR (Gy/s) | Integrated Charge / pulse (C) | Uncertainty (x10-7 C)
0.5 | 0.61986 | 4.21% | 185.9566 | 1.0043e-07 | 0.7697%
1 | 1.1351 | 4.05% | 340.531 | 2.068e-07 | 0.2163%
2 | 2.4703 | 4.501% | 741.0978 | 4.1826e-07 | 0.09148%
3 | 3.765 | 2.522% | 1129.4876 | 6.3128e-07 | 0.1497%
4 | 4.7554 | 4.328% | 1426.6098 | 8.4536e-07 | 0.2066%
5 | 6.3854 | 4.489% | 1915.6333 | 1.0584e-06 | 0.1134%
6 | 7.4054 | 2.507% | 2221.614 | 1.2442e-06 | 0.06015%

## 2 cm
FLASH DR | [5 empty] | Conventional DR
Scattering foil: | 100 | um | (0.5 us pulse) |  |  | Scattering foil: | 500 | um | (0.5 us pulse)
FWHM | 2.2 | cm |  |  |  | FWHM | 2.2 | cm
dmax | 0.5 | cm |  |  |  | dmax | 0.5 | cm
d90 | 1.62 | cm |  |  |  | d90 | 1.38 | cm
d80 | 1.99 | cm |  |  |  | d80 | 1.71 | cm
d50 | 2.85 | cm |  |  |  | d50 | 2.49 | cm
Rp | 4.34 | cm |  |  |  | Rp | 3.77 | cm
Max Sample size | 1.3 | cm | < 3% variation across field |  |  | Max Sample size | 1.3 | cm | < 4% variation across field
| 1.7 | cm | < 8% variation across field |  |  |  | 1.7 | cm | < 11% variation across field
pulse width | Dose / pulse (Gy) | Uncertainty (%) | Average DR (Gy/s) | Integrated Charge / pulse (C) | Uncertainty (%) | pulse width | Dose / pulse (Gy) | Uncertainty (Gy) | Average DR (Gy/s) | Integrated Charge / pulse (C) | Uncertainty (%)
0.5 | 5.9138 | 3.964% | 1774.1452 | 1.0043e-07 | 0.7697% | 0.5 | 0.089316 | 0.00021704 | 0.89316 | 9.3034e-09 | 3.822%
1 | 10.6914 | 3.839% | 3207.4202 | 2.068e-07 | 0.2163%
2 | 23.1913 | 4.383% | 6957.3921 | 4.1826e-07 | 0.09148%
3 | 35.2879 | 2.41% | 10586.3844 | 6.3128e-07 | 0.1497%
4 | 44.3739 | 4.308% | 13312.1613 | 8.4536e-07 | 0.2066%
5 | 59.386 | 4.5% | 17815.7932 | 1.0584e-06 | 0.1134%
6 | 68.6308 | 2.559% | 20589.2403 | 1.2442e-06 | 0.06015%
Scattering foil: | 200 | um | (0.5 us pulse)
FWHM | 2.2 | cm
dmax | 0.5 | cm
d90 | 1.52 | cm
d80 | 1.89 | cm
d50 | 2.76 | cm
Rp | 4.25 | cm
Max Sample size | 1.3 | cm | < 3% variation across field
| 1.7 | cm | < 8% variation across field
pulse width | Dose / pulse (Gy) | Uncertainty (%) | Average DR (Gy/s) | Integrated Charge / pulse (C) | Uncertainty (%)
0.5 | 3.0869 | 3.964% | 926.0683 | 1.0098 | 0.0047183
1 | 5.6884 | 3.839% | 1706.5344 | 2.0998 | 0.0073664
2 | 11.617 | 4.383% | 3485.0936 | 4.2263 | 0.0099058
3 | 18.0134 | 2.41% | 5404.0227 | 6.335 | 0.01116
4 | 23.3844 | 4.308% | 7015.3118 | 8.546 | 0.011181
5 | 30.175 | 4.5% | 9052.5047 | 10.7 | 0.0091493
6 | 35.7246 | 2.559% | 10717.3681 | 12.5067 | 0.0050248
Scattering foil: | 500 | um | (0.5 us pulse)
FWHM | 2.2 | cm
dmax | 0.5 | cm
d90 | 1.4 | cm
d80 | 1.73 | cm
d50 | 2.55 | cm
Rp | 3.94 | cm
Max Sample size | 1.3 | cm | < 2% variation across field
| 1.7 | cm | < 8% variation across field
pulse width | Dose / pulse (Gy) | Uncertainty (%) | Average DR (Gy/s) | Integrated Charge / pulse (C) | Uncertainty (%)
0.5 | 1.041 | 3.964% | 312.2922 | 1.0098 | 0.0047183
1 | 1.7571 | 3.839% | 527.1159 | 2.0998 | 0.0073664
2 | 3.8647 | 4.383% | 1159.4241 | 4.2263 | 0.0099058
3 | 6.0329 | 2.41% | 1809.8741 | 6.335 | 0.01116
4 | 7.3213 | 4.308% | 2196.3877 | 8.546 | 0.011181
5 | 10.0066 | 4.5% | 3001.9897 | 10.7 | 0.0091493
6 | 12.0312 | 2.559% | 3609.3691 | 12.5067 | 0.0050248

## 5 cm (2) (hidden)
FLASH DR | [6 empty] | Conventional DR
Scattering foil: | 200 | um | (0.5 us pulse) |  |  |  | Scattering foil: | 500 | um | (0.5 us pulse)
FWHM | 5.5 | cm |  |  |  |  | FWHM | 5.5 | cm
dmax | 0.6 | cm |  |  |  |  | dmax | 0 | cm
d90 | 1.98 | cm |  |  |  |  | d90 | 1.51 | cm
d80 | 2.52 | cm |  |  |  |  | d80 | 1.97 | cm
d50 | 3.46 | cm |  |  |  |  | d50 | 2.86 | cm
Rp | 4.79 | cm |  |  |  |  | Rp | 4.11 | cm
Max Sample size | 3.3 | cm | < 3% variation across field |  |  |  | Max Sample size | 3.3 | cm |  | < 2% variation across field
| 4.4 | cm | < 8% variation across field |  |  |  |  | 4.4 | cm |  | < 6% variation across field
pulse width | Dose / pulse (Gy) | Uncertainty (Gy) | Average DR (Gy/s) | integrated charge / pulse (x10-7 C) | Uncertainty (x10-7 C) |  | pulse width | Dose / pulse (Gy) | Uncertainty (Gy) | Average DR (Gy/s) | integrated charge / pulse (C) (x10-9) | Uncertainty (x10-9 C)
0.5 | 2.7953 | 0.24307 | 838.5938 | 1.0098 | 0.0047183 |  | 0.5 | 0.08633 | 0.0013383 | 0.8633 | 9.2911 | 0.362
1 | 5.5906 | 0.48614 | 1677.1876 | 2.0998 | 0.0073664
2 | 11.1813 | 0.97228 | 3354.3752 | 4.2263 | 0.0099058
3 | 16.7719 | 1.4584 | 5031.5628 | 6.335 | 0.01116
4 | 22.3625 | 1.9446 | 6708.7504 | 8.546 | 0.011181
5 | 27.9531 | 2.4307 | 8385.938 | 10.7 | 0.0091493
6 | 33.5438 | 2.9168 | 10063.1257 | 12.5067 | 0.0050248 | 0.086957
Scattering foil: | 500 | um | (0.5 us pulse)
FWHM | 5.5 | cm
dmax | 0 | cm
d90 | 1.63 | cm
d80 | 2.12 | cm
d50 | 3.05 | cm
Rp | 4.35 | cm
Max Sample size | 3.3 | cm | < 2% variation across field
| 4.4 | cm | < 6% variation across field
pulse width | Dose / pulse (Gy) | Uncertainty (Gy) | Average DR (Gy/s) | integrated charge / pulse (x10-7 C) | Uncertainty (x10-7 C)
0.5 | 1.0796 | 0.12469 | 323.885 | 1.0098 | 0.0047183
1 | 2.1592 | 0.24937 | 647.77 | 2.0998 | 0.0073664
2 | 4.3185 | 0.49875 | 1295.5401 | 4.2263 | 0.0099058
3 | 6.4777 | 0.74812 | 1943.3101 | 6.335 | 0.01116
4 | 8.6369 | 0.9975 | 2591.0802 | 8.546 | 0.011181
5 | 10.7962 | 1.2469 | 3238.8502 | 10.7 | 0.0091493
6 | 12.9554 | 1.4962 | 3886.6202 | 12.5067 | 0.0050248

## 5 cm
FLASH DR | [5 empty] | Conventional DR
Scattering foil: | 200 | um | (0.5 us pulse) |  |  | Scattering foil: | 500 | um | (0.5 us pulse)
FWHM | 5.5 | cm |  |  |  | FWHM | 5.5 | cm
dmax | 0.6 | cm |  |  |  | dmax | 0 | cm
d90 | 1.98 | cm |  |  |  | d90 | 1.51 | cm
d80 | 2.52 | cm |  |  |  | d80 | 1.97 | cm
d50 | 3.46 | cm |  |  |  | d50 | 2.86 | cm
Rp | 4.79 | cm |  |  |  | Rp | 4.11 | cm
Max Sample size | 3.3 | cm | < 3% variation across field |  |  | Max Sample size | 3.3 | cm |  | < 2% variation across field
| 4.4 | cm | < 8% variation across field |  |  |  | 4.4 | cm |  | < 6% variation across field
pulse width | Dose / pulse (Gy) | Uncertainty (%) | Average DR (Gy/s) | Integrated Charge / pulse (C) | Uncertainty (%) | pulse width | Dose / pulse (Gy) | Uncertainty (%) | Average DR (Gy/s) | Integrated Charge / pulse (C) | Uncertainty (%)
0.5 | 2.3617 | 3.964% | 708.504 | 1.0043e-07 | 0.7697% | 0.5 | 0.08633 | 0.0013383 | 0.8633 | 9.3034e-09 | 3.822%
1 | 4.3301 | 3.839% | 1299.0336 | 2.068e-07 | 0.2163%
2 | 9.4404 | 4.383% | 2832.1244 | 4.1826e-07 | 0.09148%
3 | 14.4103 | 2.41% | 4323.0937 | 6.3128e-07 | 0.1497%
4 | 18.185 | 4.308% | 5455.4932 | 8.4536e-07 | 0.2066%
5 | 24.4182 | 4.5% | 7325.4539 | 1.0584e-06 | 0.1134%
6 | 28.3629 | 2.559% | 8508.8639 | 1.2442e-06 | 0.06015%
Scattering foil: | 500 | um | (0.5 us pulse)
FWHM | 5.5 | cm
dmax | 0 | cm
d90 | 1.63 | cm
d80 | 2.12 | cm
d50 | 3.05 | cm
Rp | 4.35 | cm
Max Sample size | 3.3 | cm | < 2% variation across field
| 4.4 | cm | < 6% variation across field
pulse width | Dose / pulse (Gy) | Uncertainty (%) | Average DR (Gy/s) | Integrated Charge / pulse (C) | Uncertainty (%)
0.5 | 1.0086 | 3.964% | 302.5725 | 1.0043e-07 | 0.7697%
1 | 1.8377 | 3.839% | 551.304 | 2.068e-07 | 0.2163%
2 | 4.0021 | 4.383% | 1200.629 | 4.1826e-07 | 0.09148%
3 | 6.108 | 2.41% | 1832.3906 | 6.3128e-07 | 0.1497%
4 | 7.7457 | 4.308% | 2323.7154 | 8.4536e-07 | 0.2066%
5 | 10.3858 | 4.5% | 3115.7542 | 1.0584e-06 | 0.1134%
6 | 12.0781 | 2.559% | 3623.4324 | 1.2442e-06 | 0.06015%

## 6 cm (2) (hidden)
FLASH DR | [5 empty] | Conventional DR
Scattering foil: | 200 | um | (0.5 us pulse) |  |  | Scattering foil: | 500 | um | (0.5 us pulse)
FWHM | 6.5 | cm |  |  |  | FWHM | 6.5 | cm
dmax | 0.5 | cm |  |  |  | dmax | 0 | cm
d90 | 0.95 | cm |  |  |  | d90 | 1.35 | cm
d80 | 2.54 | cm |  |  |  | d80 | 1.84 | cm
d50 | 3.54 | cm |  |  |  | d50 | 2.81 | cm
Rp | 4.87 | cm |  |  |  | Rp | 4.2 | cm
[6 empty] | 500 | um
Max Sample size | 3.9 | cm | < 3% variation across field |  |  | Max Sample size | 3.9 | cm | < 1% variation across field
| 5.2 | cm | < 7% variation across field |  |  |  | 5.2 | cm | < 3% variation across field
pulse width | Dose / pulse (Gy) | Uncertainty (Gy) | Average DR (Gy/s) | integrated charge / pulse (x10-7 C) | Uncertainty (x10-7 C) | pulse width | Dose / pulse (Gy) | Uncertainty (Gy) | Average DR (Gy/s) | integrated charge / pulse (C) (x10-9) | Uncertainty (x10-9 C)
0.5 | 2.3418 | 0.14119 | 702.5538 | 1.0098 | 0.0047183 | 0.5 | 0.08399 | 0.00018109 | 0.8399 | 9.2911 | 0.362
1 | 4.6837 | 0.28238 | 1405.1075 | 2.0998 | 0.0073664
2 | 9.3674 | 0.56476 | 2810.215 | 4.2263 | 0.0099058
3 | 14.0511 | 0.84715 | 4215.3225 | 6.335 | 0.01116
4 | 18.7348 | 1.1295 | 5620.43 | 8.546 | 0.011181
5 | 23.4185 | 1.4119 | 7025.5375 | 10.7 | 0.0091493
6 | 28.1022 | 1.6943 | 8430.645 | 12.5067 | 0.0050248
Scattering foil: | 500 | um | (0.5 us pulse)
FWHM | 6.5 | cm
dmax | 0 | cm
d90 | 1.42 | cm
d80 | 1.98 | cm
d50 | 3.01 | cm
Rp | 4.4 | cm
Max Sample size | 3.9 | cm | < 2% variation across field
| 5.2 | cm | < 2% variation across field
pulse width | Dose / pulse (Gy) | Uncertainty (Gy) | Average DR (Gy/s) | integrated charge / pulse (x10-7 C) | Uncertainty (x10-7 C)
0.5 | 1.0596 | 0.10993 | 317.8682 | 1.0098 | 0.0047183
1 | 2.1191 | 0.21985 | 635.7364 | 2.0998 | 0.0073664
2 | 4.2382 | 0.4397 | 1271.4727 | 4.2263 | 0.0099058
3 | 6.3574 | 0.65956 | 1907.2091 | 6.335 | 0.01116
4 | 8.4765 | 0.87941 | 2542.9455 | 8.546 | 0.011181
5 | 10.5956 | 1.0993 | 3178.6818 | 10.7 | 0.0091493
6 | 12.7147 | 1.3191 | 3814.4182 | 12.5067 | 0.0050248

## 6 cm
FLASH DR | [5 empty] | Conventional DR
Scattering foil: | 200 | um | (0.5 us pulse) |  |  | Scattering foil: | 500 | um | (0.5 us pulse)
FWHM | 6.5 | cm |  |  |  | FWHM | 6.5 | cm
dmax | 0.5 | cm |  |  |  | dmax | 0 | cm
d90 | 0.95 | cm |  |  |  | d90 | 1.35 | cm
d80 | 2.54 | cm |  |  |  | d80 | 1.84 | cm
d50 | 3.54 | cm |  |  |  | d50 | 2.81 | cm
Rp | 4.87 | cm |  |  |  | Rp | 4.2 | cm
[6 empty] | 500 | um
Max Sample size | 3.9 | cm | < 3% variation across field |  |  | Max Sample size | 3.9 | cm | < 1% variation across field
| 5.2 | cm | < 7% variation across field |  |  |  | 5.2 | cm | < 3% variation across field
pulse width | Dose / pulse (Gy) | Uncertainty (%) | Average DR (Gy/s) | Integrated Charge / pulse (C) | Uncertainty (%) | pulse width | Dose / pulse (Gy) | Uncertainty (%) | Average DR (Gy/s) | Integrated Charge / pulse (C) | Uncertainty (%)
0.5 | 2.1781 | 2.91% | 653.4281 | 1.0043e-07 | 0.7697% | 0.5 | 0.08399 | 0.00018109 | 0.8399 | 9.3034e-09 | 3.822%
1 | 3.9907 | 2.862% | 1197.1976 | 2.068e-07 | 0.2163%
2 | 8.6597 | 3.735% | 2597.9184 | 4.1826e-07 | 0.09148%
3 | 13.2306 | 1.304% | 3969.1682 | 6.3128e-07 | 0.1497%
4 | 16.6955 | 3.934% | 5008.6585 | 8.4536e-07 | 0.2066%
5 | 22.4304 | 4.249% | 6729.1318 | 1.0584e-06 | 0.1134%
6 | 26.0457 | 2.26% | 7813.7121 | 1.2442e-06 | 0.06015%
Scattering foil: | 500 | um | (0.5 us pulse)
FWHM | 6.5 | cm
dmax | 0 | cm
d90 | 1.42 | cm
d80 | 1.98 | cm
d50 | 3.01 | cm
Rp | 4.4 | cm
Max Sample size | 3.9 | cm | < 2% variation across field
| 5.2 | cm | < 2% variation across field
pulse width | Dose / pulse (Gy) | Uncertainty (%) | Average DR (Gy/s) | Integrated Charge / pulse (C) | Uncertainty (%)
0.5 | 0.96766 | 2.91% | 290.2975 | 1.0043e-07 | 0.7697%
1 | 1.766 | 2.862% | 529.7889 | 2.068e-07 | 0.2163%
2 | 3.8399 | 3.735% | 1151.9557 | 4.1826e-07 | 0.09148%
3 | 5.8821 | 1.304% | 1764.6445 | 6.3128e-07 | 0.1497%
4 | 7.4471 | 3.934% | 2234.1311 | 8.4536e-07 | 0.2066%
5 | 9.9936 | 4.249% | 2998.0784 | 1.0584e-06 | 0.1134%
6 | 11.6168 | 2.26% | 3485.0397 | 1.2442e-06 | 0.06015%

## 10 cm
FLASH DR | [5 empty] | Conventional DR
Scattering foil: | 500 | um | (0.5 us pulse) |  |  | Scattering foil: | 500 | um | (0.5 us pulse)
FWHM | 10.8 | cm |  |  |  | FWHM | 10.7 | cm
dmax | 0 | cm |  |  |  | dmax | 0 | cm
d90 | 1.31 | cm |  |  |  | d90 | 1.48 | cm
d80 | 2.18 | cm |  |  |  | d80 | 2.16 | cm
d50 | 3.2 | cm |  |  |  | d50 | 3.08 | cm
Rp | 4.44 | cm |  |  |  | Rp | 4.2 | cm
500 | um
Max Sample size | 6.4 | cm | < 3% variation across field |  |  | Max Sample size | 6.5 | cm | < 2% variation across field
| 8.6 | cm | < 6% variation across field |  |  |  | 8.6 | cm | < 3% variation across field
pulse width | Dose / pulse (Gy) | Uncertainty (%) | Average DR (Gy/s) | Integrated Charge / pulse (C) | Uncertainty (%) | pulse width | Dose / pulse (Gy) | Uncertainty (%) | Average DR (Gy/s) | Integrated Charge / pulse (C) | Uncertainty (%)
0.5 | 0.73477 | 2.635% | 220.4324 | 1.0043e-07 | 0.7697% | 0.5 | 0.057868 | 0.00019988 | 0.57868 | 9.3034e-09 | 3.822%
1 | 1.344 | 2.627% | 403.2027 | 2.068e-07 | 0.2163%
2 | 2.9272 | 3.64% | 878.1557 | 4.1826e-07 | 0.09148%
3 | 4.4633 | 1.314% | 1338.9965 | 6.3128e-07 | 0.1497%
4 | 5.6366 | 4.046% | 1690.9677 | 8.4536e-07 | 0.2066%
5 | 7.5623 | 4.467% | 2268.6806 | 1.0584e-06 | 0.1134%
6 | 8.7712 | 2.856% | 2631.3556 | 1.2442e-06 | 0.06015%

## 10 cm (2) (hidden)
FLASH DR | [5 empty] | Conventional DR
Scattering foil: | 500 | um | (0.5 us pulse) |  |  | Scattering foil: | 500 | um | (0.5 us pulse)
FWHM | 10.8 | cm |  |  |  | FWHM | 10.7 | cm
dmax | 0 | cm |  |  |  | dmax | 0 | cm
d90 | 1.31 | cm |  |  |  | d90 | 1.48 | cm
d80 | 2.18 | cm |  |  |  | d80 | 2.16 | cm
d50 | 3.2 | cm |  |  |  | d50 | 3.08 | cm
Rp | 4.44 | cm |  |  |  | Rp | 4.2 | cm
500 | um
Max Sample size | 6.4 | cm | < 3% variation across field |  |  | Max Sample size | 6.5 | cm | < 2% variation across field
| 8.6 | cm | < 6% variation across field |  |  |  | 8.6 | cm | < 3% variation across field
pulse width | Dose / pulse (Gy) | Uncertainty (Gy) | Average DR (Gy/s) | integrated charge / pulse (x10-7 C) | Uncertainty (x10-7 C) | pulse width | Dose / pulse (Gy) | Uncertainty (Gy) | Average DR (Gy/s) | integrated charge / pulse (C) (x10-9) | Uncertainty (x10-9 C)
0.5 | 0.87053 | 0.15959 | 261.1581 | 1.0098 | 0.0047183 | 0.5 | 0.057868 | 0.00019988 | 0.57868 | 9.2911 | 0.362
1 | 1.7411 | 0.31919 | 522.3162 | 2.0998 | 0.0073664
2 | 3.4821 | 0.63837 | 1044.6323 | 4.2263 | 0.0099058
3 | 5.2232 | 0.95756 | 1566.9485 | 6.335 | 0.01116
4 | 6.9642 | 1.2767 | 2089.2647 | 8.546 | 0.011181
5 | 8.7053 | 1.5959 | 2611.5809 | 10.7 | 0.0091493
6 | 10.4463 | 1.9151 | 3133.897 | 12.5067 | 0.0050248

## 2 cm (2) (hidden)
FLASH DR | [5 empty] | Conventional DR
Scattering foil: | 100 | um | (0.5 us pulse) |  |  | Scattering foil: | 500 | um | (0.5 us pulse)
FWHM | 2.2 | cm |  |  |  | FWHM | 2.2 | cm
dmax | 0.5 | cm |  |  |  | dmax | 0.5 | cm
d90 | 1.62 | cm |  |  |  | d90 | 1.38 | cm
d80 | 1.99 | cm |  |  |  | d80 | 1.71 | cm
d50 | 2.85 | cm |  |  |  | d50 | 2.49 | cm
Rp | 4.34 | cm |  |  |  | Rp | 3.77 | cm
Max Sample size | 1.3 | cm | < 3% variation across field |  |  | Max Sample size | 1.3 | cm | < 4% variation across field
| 1.7 | cm | < 8% variation across field |  |  |  | 1.7 | cm | < 11% variation across field
pulse width | Dose / pulse (Gy) | Uncertainty (Gy) | Average DR (Gy/s) | integrated charge / pulse (x10-7 C) | Uncertainty (x10-7 C) | pulse width | Dose / pulse (Gy) | Uncertainty (Gy) | Average DR (Gy/s) | integrated charge / pulse (C) (x10-9) | Uncertainty (x10-9 C)
0.5 | 6.1501 | 0.75872 | 1845.0194 | 1.0098 | 0.0047183 | 0.5 | 0.089316 | 0.00021704 | 0.89316 | 9.2911 | 0.362
1 | 12.3001 | 1.5174 | 3690.0387 | 2.0998 | 0.0073664
2 | 24.6003 | 3.0349 | 7380.0774 | 4.2263 | 0.0099058
3 | 36.9004 | 4.5523 | 11070.1161 | 6.335 | 0.01116
4 | 49.2005 | 6.0698 | 14760.1548 | 8.546 | 0.011181
5 | 61.5006 | 7.5872 | 18450.1935 | 10.7 | 0.0091493
6 | 73.8008 | 9.1047 | 22140.2322 | 12.5067 | 0.0050248
Scattering foil: | 200 | um | (0.5 us pulse)
FWHM | 2.2 | cm
dmax | 0.5 | cm
d90 | 1.52 | cm
d80 | 1.89 | cm
d50 | 2.76 | cm
Rp | 4.25 | cm
Max Sample size | 1.3 | cm | < 3% variation across field
| 1.7 | cm | < 8% variation across field
pulse width | Dose / pulse (Gy) | Uncertainty (Gy) | Average DR (Gy/s) | integrated charge / pulse (x10-7 C) | Uncertainty (x10-7 C)
0.5 | 3.399 | 0.21058 | 1019.6992 | 1.0098 | 0.0047183
1 | 6.798 | 0.42117 | 2039.3984 | 2.0998 | 0.0073664
2 | 13.596 | 0.84233 | 4078.7969 | 4.2263 | 0.0099058
3 | 20.394 | 1.2635 | 6118.1953 | 6.335 | 0.01116
4 | 27.192 | 1.6847 | 8157.5938 | 8.546 | 0.011181
5 | 33.99 | 2.1058 | 10196.9922 | 10.7 | 0.0091493
6 | 40.788 | 2.527 | 12236.3907 | 12.5067 | 0.0050248
Scattering foil: | 500 | um | (0.5 us pulse)
FWHM | 2.2 | cm
dmax | 0.5 | cm
d90 | 1.4 | cm
d80 | 1.73 | cm
d50 | 2.55 | cm
Rp | 3.94 | cm
Max Sample size | 1.3 | cm | < 2% variation across field
| 1.7 | cm | < 8% variation across field
pulse width | Dose / pulse (Gy) | Uncertainty (Gy) | Average DR (Gy/s) | integrated charge / pulse (x10-7 C) | Uncertainty (x10-7 C)
0.5 | 1.1137 | 0.17421 | 334.1158 | 1.0098 | 0.0047183
1 | 2.2274 | 0.34843 | 668.2317 | 2.0998 | 0.0073664
2 | 4.4549 | 0.69686 | 1336.4634 | 4.2263 | 0.0099058
3 | 6.6823 | 1.0453 | 2004.6951 | 6.335 | 0.01116
4 | 8.9098 | 1.3937 | 2672.9268 | 8.546 | 0.011181
5 | 11.1372 | 1.7421 | 3341.1585 | 10.7 | 0.0091493
6 | 13.3646 | 2.0906 | 4009.3902 | 12.5067 | 0.0050248